## Sheet1
start: 01/24/2017 13:00
Step 0: Clean out sensor chambers with condensed air spray
Step 1: Wipe down interior srufaces of box with kimwipes and alcohol; wipe underside of each Nest (white plastic side) with kimwipes and alcohol
Step 2: Spot check operation of each nest visually (red LEDs blinking)
Step 3: Change onboard clocks to match between devices
Step 4: Check operation by SSH-ing in (checking that both the C and python scripts are logging)
Step 5: Ziptie box shut
Step 6: Begin pumping clean air into calibration box at 5.0446 LPM (Measured with MesaLabs Defender 510, Rev C1, S/N 136834)
Step 7: Remove pump, and attach DustTrak II 8530 (HECHH Monitor "X1")… DT loaded with a blank filter, impactor cleaned and oiled, flow rate confirmed with Defender510 at 2.9965 LPM, and zero calibrated. DT Runmode 4 hours, 1 sec log interval, 1 second time constant, 1x PD calibration coefficient
Step 8: Remove HEPA and DustTrak, filliung inlet and outlet with rubber stoppers and capping with rubber caps
Step 9: Recheck Nest Function via SSH
Step 10: begin incubator testing
Step 11: end incubator testing, allow several minutes for box to come back to within 8C of room temp / temp at which zero cal occurred (as per DT manual… temp during zero cal was ~ 20C) (Ice bpacks were placed atop the box to assist in the effort)
Step 12: Measure post-test concentration inside box (unseal one end-- attach end immediately to DustTrak then sample for 10 minutes)
[8 empty] | DT File: M annual_018 | [7 empty] | DT File Manual_019 |  |  |  |  | DT File Manual_020
[6 empty] | Pulling Clean Air Into Box |  | ZERO CAL, Box Concentration, pre-testing |  |  | Room Concentration, pre-testing |  | Box concentration, pre-testing |  |  | Incubator concentration, mid-testing |  |  |  |  | Box concentration, post-testing |  |  | Room concentration, post-testing
Device | Visual Function Check | SSH Templog Check | SSH ADC log check | Check Notes | Box Pump LPM (SKC Model 224 PCXR8) | Begin Box Pump | End Box Pump | Box Reading DustTrak File Name | Box DustTrak Start (with Zero Filter on) | End DT With Zero Filter | Begin Background DT Reading (Room Value) | End Background DT Reading (Room Value) | Begin Box DT Reading | End Box DT Reading | Check Notes | Incubator Background Dusttrak Filename | DustTrak On with zero filter | Begin Incubator Background Reading (20C) | End Incubator Backgroudn Reading (20C) | Notes | Post-Testing, Box Concentration Measurement DustTrak File Name | Start (w/plug in box inlet) | End | Begin Background DT Reading (Room Value) | End Background DT Reading (Room Value) | Put HEPA on DustTrak | End
alpha | 1 | 1 | 1 | Clocks synced within a second | 5.0446 | 13:35 | 14:12 | Manual_018 | 14:12 | 14:15 | 14:16 | 14:26 | 14:27 | 14:37 | Clocks synced within a second | Manual_019 | 16:56 | 16:59 | 17:14 | Vibration during incubator DT testing may have caused some interesting vacillation with the Nest sensors… likely not | Manual_020 | 20:41 | 20:51 | 20:52 | 21:02 | 21:03 | 21:13
beta | 1 | 1 | 1 | Clocks synced within a second | [10 empty] | Clocks synced within a second
epsilon | 1 | 1 | 1 | Clocks synced within a second | [10 empty] | Clocks synced within a second
Note: Box is 10 gallons = 37.85 Liters… 8 minutes should be enough time to clear the box at 5 LPM
Time | Event (temps are Incubator temps) | Notes
14:40 | Box placed in Incubator, set to 0C
14:45 | arrived at 0 C
15:40 | Set to 10 C
16:40 | Set to 20C
17:40 | Set 30 C
18:44 | Set 40 C
19:40 | Set 48 C | This is the max temp setting
20:10 | End 48 C, remove from incubator | According to Rpi, the Rpi 3 begins to throttle speed at 80C… look into effects of this on data
20:30 | SHT sensors reading ~ 26C - 27C inside box
20:52 | Lid Removed from box
21:03:45 AM | Power Off Nests (given 25 seconds to boot off via ssh command, followed by hard shutoff)
Note: PM2.5 impactor may not have been fully sealed, though it was tightly fastened-- probably not a major deal