## Sheet1
DustTrak Specs
|  |  | Device | Filter ID | File ID | Calibrated Flow (Defender 510,S/N 136834) | Datetime Set | Zero Cal | Log Interval | Time Constant | Sample Length Max | Note 1 | Note 2 | Nopte 3
|  |  | DT X1, S/N 8530124209 | NEST_08 | MANUAL_028 | 2998.5 ml/min (taken without impactor) | x | x | 1 sec | 1 sec | 7 hrs | Impactor not applied; OK because air was at zero PM, but double check file | Light dump and impactor cleaned prior to test
|  |  | DT Y, S/N. 8530124211 | NEST_09 | MANUAL_005 | 3005.6 ml/mIN (taken without impactor) | x | x | 1 sec | 1 sec | 7 hrs | Impactor not applied; OK because air was at zero PM, but double check file | Light dump and impactor cleaned prior to test | Linty-hair found on filter after the end of sample and removed. It was relatively stuck into the filter.
|  |  | DT X1 | NEST_08 | MANUAL_029 | 2983.6 (taken post-test, with impactor) | x | x | 1 sec | 1 sec | 7 hrs
|  |  | DT Y | NEST_09 | MANUAL_006 | 2970.5 (taken post-test, with impactor) | x | x | 1 sec | 1 sec | 7 hrs
|  |  | DT X1 | NA | MANUAL_030 |  | x | x | 1 sec | 1 sec | 7 hrs
|  |  | DT Y | NA | MANUAL_007 |  | x | x | 1 sec | 1 sec | 7 hrs
Time | Event |  | Notes
9:49 | Rpi Clocks Synced to the internet and each other. Verified. |  | Smoke produced primarily by applying hot metal (solder iron) to wood shingles
9:50 | Chamber sealed, HEPA filter / pull-through enabled (begin approaching zero) |  | Some smoke early on produced by burning wood shingles or wood sticks with kerosene lighter ( ~ 1 minute worth).
13:02 | Begin Zero sample - baseline
13:04 | Chamber fan turned on (was accidentally off at start)
13:06 | Baseline Zero paused to apply PM2.5 impactors… First DT Files stopped
13:17 | Baseline Zero Resumed… Second DT Files begun
13:48 | Raise Concentration to 70 mg/m3 (according to DT screen)
14:18 | Back to Zero
14:48 | Raise Concentration to 5 mg/m3 (according to DT screen) |  | <- had some trouble reaching 5 mg/m3. Took ~ 10 minutes to get there
15:18 | Back to zero
15:48 | Raise Concentration to 0.5 mg/m3 |  | <- increase to 45 minutes to account for decrease in sample rate
16:33 | Back to zero
17:03 | Raise Concentration to 0.100 mg/m3 |  | <- 45 minutes to account for decrease in sample rate
17:43 | Back to Zero
18:08 | Raise Concentration to 10 mg/m3
18:38 | Slow descent back to zero
19:00 | Reached zero
20:00 | Finished with testing (DT Tubes detached from chamber, tubes not sealed)
20:00 | Zero filter applied to each DustTrak
20:06 | Zero filters removed
TEST FOR RESPONSIVENESS - NO FILTER IN DUSTTRAK (empty cassettes w/steel mesh backing pads)
20:18 | Start at zero - Fan On
20:20 | Fan Off
20:23 | Raise Concentration to (uncalibrated) 10mg/m3
20:27 | Chamber sealed (except for DT outlets)
20:33 | Dropped to 5 mg/m3 by opening valve and applying HEPA filter to lid hole
20:35 | Chamber resealed
20:42 | Fan on | 60 seconds
20:43 | Fan Off
20:45 | Dropped to 1 mg/m3 by opening valve and applying HEPA filter to lid hole
20:47 | Chamber resealed
20:54 | Fan On | 60 seconds
20:55 | Fan Off
20:57 | Zero by opening valve and applying HEPA filter to lid hole
21:45 | Far vent outlet closed; HEPA on lid hole remained, and DT continued pumping at a combined 3 LPM
21:23 | Finish